## Sheet1
|  |  | Uncertainty squared |  |  | meas u^2 |  | comp u^2 | Notes | Interpretation | Measure by
|  | Spec max +/- | XYZ (meters^2) | RxRyRz (radians^2) | Measured +/- | XYZ (meters^2) | Comp +/- | XYZ (meters^2) |  | Total indicator reading / 2 | Measurement from pole top will include moment effects
Z axis | Parker Daedal 806004CT
| Positional accuracy (um) | 15 | 2.25e-10 |  | 65 | 4.225e-09 | 16.25 | 2.6406e-10 |  |  | Gauge blocks
| Straight line (um) | 8 | 6.4e-11 |  | 110 | 1.21e-08 | 27.5 | 7.5625e-10 |  | Off axis lateral? | Indicator
| Flatness (um) | 8 | 6.4e-11 |  | 20 | 4e-10 | 20 | 4e-10 |  | Off axis vertical? | Indicator
| Repeatability (um) | 5 | [8 empty] | Indicator
XY axis | Parker Daedal 300 AT series, 310062AT
| Repeatability, bidirectional (um) | 12 | [8 empty] | Indicator
| X-Y squareness (urad) | 291 | 2.117e-10 |  | 0 | 0 | 0 | 0 | Moment is +/- 50mm XY travel | in meas case is absorbed into straight and flat | Square
| Positional accuracy (um) | 24 | 5.76e-10 |  | 191 | 3.6481e-08 | 47.75 | 2.2801e-09 |  |  | Gauge blocks
| Straightness and flatness (um) | 76 | 5.776e-09 |  | 14 | 1.96e-10 | 28 | 7.84e-10 |  | Interpreting as XY vector error | Indicator
Rz axis | Parker Daedal 200RT series, 20801RTES
| Repeatability, unidirectional (urad) | 145 | [8 empty] | Indicator over moment
| Accuracy (urad) | 2900 |  | 8.41e-06 |  |  |  |  | Ouch! |  | Use square and test X move deviation at 30, 60, 90
| Runout (um) | 75 | 5.625e-09 | [6 empty] | XY deviation with no Z moment | Runout in Rz alignment test
| Concentricity (um) | 127 | [7 empty] | Rz center vs stage body? | Ignore: doesn't matter
| Y off-axis non-repeatability (um) |  |  |  | 15 | [5 empty] | Runout in Rz alignment test
| Wobble (urad) | 291 | 2.117e-10 | 8.4681e-08 |  |  |  |  | Moment is +/- 50mm Z travel | Off-axis RxRy? Includes non-square of Rz vs base? | Change in Rz runout vs Z height
Source fixture | Concentricity (um) |  | 1.69e-10 |  | 13 | [5 empty] | Fixture ball deviation across positions
| Squareness (urad) |  | 4e-10 | 1e-08 | 100 |  |  |  | Moment is 200mm sensor distance |  | Rocking test vs. square
Sensor fixture | Concentricity (um) |  | 4e-10 |  | 20 | [5 empty] | Fixture ball deviation across positions
| Squareness (urad) |  |  | 9e-08 | 300
| Repeatability (um) |  | 1.69e-10 |  | 13 | [5 empty] | Variation of Y indication by refixturing (includes Y move away and back)
|  |  | 117.8618 | 2931.6686 |  | 245.7167 |  | 107.0471
|  |  | XYZ (um) | RxRyRz (urad) |  | XYZ (um) |  | XYZ (um)
|  |  | Spec uncertainty |  |  | Measured u |  | Compensated